## Slide 1: 2nd NIST Workshop on Cement Materials Characterization
Don Broton
CTLGroup
Gaithersburg, MD July, 2017
XRF ANALYSIS

## Slide 2: X-ray Analysis
XRF is elemental composition, fast, accurate
Perfect for production
Alternate C114 wet methods too time consuming
We report as the oxide by convention
We can run solids or liquids
Metals reported as the element

## Slide 3: Schematic Diagram

## Slide 4: Steradian

## Slide 5: X-ray tube spectrum
Electrons accelerated towards target
hits target and gives up energy step wise which gives rise to the continuum
More Kv shifts short wavelength limit to the left thus exciting higher atomic number elements
Higher energy/shorter wavelengths
Lower energy/longer wavelengths

## Slide 6: Ionization

## Slide 7: Emmision of characteristic x-ray spectra

## Slide 8: Spectrometer Configuration

## Slide 9: Infinite Thickness
In a typical portland cement briquette (35.5% porosity) the mass of the specimen actually analyzed is about 2.5 mg for sodium, 50mg for calcium!

## Slide 10: Why Consider Sampling?
Analytical chemists use samples; they are in charge of QC and usually blamed when analytical results, which depend heavily on sampling errors, reveal discrepancies. Pierre M. Guy, “The analaytical and economic correctness in sampling” Anal Chim Acta 190, 13-23, 1986
Notes: I you consider the amount of time it takes to properly assess the sample & specimen prepapation protcol for any sample stream it is extremely small to the amount of time it will take to determine an error has been made, find the source, and then start over. I can recal not too long ago when a researched was short on time for a project so he grabbed one rock from a barrel of aggregate. Processed the one rock in ring and puck mill, then submitted a 1oz sample for analysis. After obtaining the results made many mixes, blending for hours at a time, clinkered the blends and then analyzed them. Surprise , surprise- the composition was nothing he expected. What a waste of an entire week of laboring! Don’t fall into that time crunch trap.

## Slide 11: Different Shapes & Sizes
What is different about this sample?
Notes: Samples come in variety of shapes and sizes. Choosing the appropriate sample/specimen technique will insure the most accurate of data . It is always important to note the condition as well. Was the sample wet, dry, macroscopically homogeneous? Notes are important because with the amount of samples you may see in a year it will be difficult if not impossible to remember all of them.

## Slide 12: Sources of Contamination
Sampling Equipment
Handling
Storage Containers
Processing / Crushing / Grinding Equipment
Notes: Any equipment can be sources of contamination. It is best to know compositions of the materials and the analyses required. One analyst came to the lab for a jarto place a liquid that needed testing. The sample was a basic solution and needed testing for sodium. Is a sodium silicate glass appropriate?

## Slide 13: Crushing / Grinding
Jaw Crusher
Disc Pulverizer
Blueler Mill
Notes: Each of these pieces of equipment can and often does contribute to contamination. Using an iron disc pulverizer is not appropriate when grinding slag. It has been done at least once that I know of- there was not much left of the mill when the sample was passed through. How much iron contamination do you think was added?
A LOT!

## Slide 14: Jaw Crusher
Notes: Jaw crushers can also contribute to contamination. All equipment is made from materials with certain hardness. When this is exceeded by the sample which do you think suffers? Right . The equipment. Inspection and knowledge of the materials that are processed are really the key to success. Do not overlook the sample prep technician. He/she is you r greatest asset! Remember contamination can be lurking in crevices too!

## Slide 15: Disc Pulverizer
Notes: The type of material selected as the grinding surface is the key. They come in a variety of materials- choose the right one. Remember that the equipment is moving and has a good deal of momentum. It is not the place for loose ties, belts or other articles of clothing. Safety glasses are a must! Most overlook the potentail hazards of noise and dust. Don’t you! It is far better to use a mask, ear and eye protection when processing. Wear it- you won’t be dissappointed!

## Slide 16: Disc Pulverizer
Notes: The alumina disc pulveriser is one way of reducing particle sizes to 100% passing a number 100 sieve. The plates should be parrelel to each other and NEVER banged closed. Chipping of the surface will result and may also contaminate with those chips falling into the sample. Inspect frequently for damage. If more than one person uses the equipment be sure to inspect each time BEFORE you operate it.

## Slide 17
Contamination ?
Notes: Clearly, knowing the composition of the processing equipment and the relative hardness compared the the sample is key in evaluating the use.

## Slide 18: Ring & Puck Mill (or any process equipment) What is it made of/elements of interest?
Notes: There are several types of mills to reduce the fineness of samples. Tingsten carbide, iron, alumina and others. WC is the most energetic and also weighs the most. It is brittle as one of my technicians can tell you first hand. A single drop on the floor results in a large chip! Floor matts reduces this chance and is a good idea in general. It reduces fatiguie induced by walking on hard floors. Safety shoes also a good idea.

## Slide 19: Ring & Puck Mill
Notes: Since there is a tremndous amount of energy involved in processing equipment don’t be caught offguard by loose fitting pieces. Covering the equipment to reduce the chances of being hit and reduces noise. If you have to force something it probably is not installed right or working properly. Investigate BEFORE an accident occurs.

## Slide 20: Contamination
Notes: The blue disc is cobalt contaminated.

## Slide 21: Sampling from a Bottle
Notes: Anytime a sub-sample is taken from a bottle it should be mixed prior opening. A figure 8 motion in a jar that is no more than 2/3rd filled is most chosen. For your benefit take a bimodal mixture of two materials and place it in a jar. Rotate several times to see how it mixes. Then tap the jar several times on the table top. Watch how the particles segregate again! What about moisture? Hygroscopic materials will adsorp at the surface. Don’t bias your result!

## Slide 22: Grinding Time vs Fineness (with different mill types)
Notes: Usually 4-5 minutes in a WC mill is effective in reducing the particle size of the sample. It is recommended that a PSD be obtained to insure the proper grinding time, mill, and grinding aid. Don’t be fooled. Each mill has a different weight, each machine different speeds. Each will affect grinding efficiency.

## Slide 23: Chemical Grinding Aids
Propylene glycol
Aspirin
Liquid freon
Soaps
Detergents
Water
Commercial Tablets-Grinder/Binder?
Notes: There are many types of grinding aids available. Each will yield results that may vary from specimen type. Trial sizes of many are usually available from the manufacturers. Remember though that anything you put into the mill will be part of your sample. Make a blank of the grinding aid itself and check for the normal elements of interest. Who would choose to put sodium in a sample when you are analyzing for that element?

## Slide 24: Particle Size Distribution
Notes: The effect of grinding time and grinding aid on the particle size reduction can easily be determined. This slide shows a normal cement without grinding and tow types of grinding aid. Note: mass mean diameter of propylene glycol ground sample is about 5 microns.

## Slide 25: Commercial Binders
Notes: Binders/grinding aids come in all shapes and sizes. If the amount of binder used in the processing changes so does the dilution and may result in biased results. Low/high analytical totals or poor precision may result.

## Slide 26: Die Assembly
Notes: Typical die assembly for pressing powders. The surface of the die that comes in contact with the specimen should be flat, clean and mirror like for the best results. Polishing must be done carefully so not to misshape the disc.

## Slide 27: Die (Assembled)

## Slide 28: Spex X-Press
Notes: Pressed should be capable of reproducing the same pressure of about least 25 tons /50,000 pounds time and time again. Place the die assembly in carefully and load slowly. Hold for a minute the release slowly. Sudden release of pressure may cause a cracked bead.

## Slide 29: Longer Grinding Times
Notes: Effects of longer grinding times can also be show by making epoxy mounts and cross sectioning as shown here. It is always best to observe the specimen in ways not normally seen routinely. Neat looking too!

## Slide 30: Appearance of Pellet
Notes: Pellets should be homogeneous in appearance. Flecks of material may be cause for concern.

## Slide 31: Compare Colors
Notes: Compare the colors of the specimen before and after grinding. It may yield clues to contamination

## Slide 32: Pressed Powder Technique
Obtain representative sample (<100 US Standard Sieve)
Weigh 7 grams & place in ring & puck mill
add 1-2 drops propylene glycol
Grind 3-4 minutes
add binder, grind additional 30 seconds
press into briquett @ 50,000 (25 tons)
Place in dessicator until analyzed
Notes: The technique in use at CTL is describe above.

## Slide 33: Borate Fusions

## Slide 34: Why Fuse?
Calibrations look GREAT!
No particle size effects!
No Mineralogical effects!
Use pure compounds to calibrate
Measure “real” inter-element effects
Less contamination from intense grinding needed for pressed powders
No powder deposition on x-ray tube
Notes: More often than not, once an analyst has experience in fusions they never want to go back to pressed powders. Calibrations do look fantastic and are accurate if done correctly. A favorite of many.

## Slide 35: Why Not?
Volatiles evaporated (time and temp dependent)
dilution means higher detection limits
somewhat more complicated
costly- machine, platinum, & flux
Notes: There are pitfalls however to any technique. Knowing them is most valuable.

## Slide 36: Typical Materials Analyzed
Limestone
Clay
Iron ore
Fly ash
Cement
Gypsum
Shale
sludge
silica fume
slag

## Slide 37: Geologicals and Industrial by-Products
Notes: As always knowing the material is instrumental in deciding which technique(fusion/pressing) is more usefull and appropriate

## Slide 38
Powder Deposition on Tube?
None!
Notes: What a benefit. Fusion cannot deposit powder on the tube. Bad specimens can crack though.

## Slide 39: Releasing Agents
NaBr
LiBr
KI
NH4I
Need to evaluate line overlaps (Br-Al)
Measure Io of releasing agent
Optimize amount necessary!
Notes: Choose the appropraite releasing aid and use only the amount necessary. Too much the glass will bead and not fill the mold. Also watch out for line overlaps ie. Br Al. Some sample types reuire no releasing aid. Some stick more than others. Keeping the moulds clean and mirror-like will yield the best results.

## Slide 40: Flux Choices
Sodium tetraborate(very hygroscopic)
Lithium tetraborate
Lithium tetraborate/metaborate mixtures
Lithium metaborate-(prep solns)
In all cases measure Loss on ignition
In all cases fuse “blank” run and check for contaminants/purity

## Slide 41
Flux Choices

## Slide 42: Optimize Conditions of Fusion
Measure temperature of molten flux
Type S thermocouple inside of flux- also measure crucible wall temperature
Measure Loss on Fusion of material to analyze. Sulfur loss good indicator.
Minimum Time/Minimum Temp

## Slide 43: Procedure
Determine L.O.I.
Use Freshly Calcined Material or Balance for L.O.I.
Weigh 1gram Sample + 5 Grams L.O.I. Corrected Flux
Add Anti-wetting Agent (LiBr?)
Fuse,Cast & Cool
Check for Disc Integrity

## Slide 44
Mixing Sample-Right/Wrong?
Notes: Any item used to blend the sample and flux should not scratch the Pt. Also a glass rod may contaminate the sample. Best to mix in another contain then transfer.

## Slide 45
Quick Check
Notes: After mixing, you can empty the mixture onto a piece of paper to verify the homogeneity of it. Some sample types (ie clay, fly ash) may scinter or clump post ignition. Passing sample through a No 100 sieve will insure shortest fusion time possible and virtually elimante inclusions.

## Slide 46
Scintering-post Ignition requires grinding

## Slide 47
Fluxer Operation

## Slide 48
Dark Blue Disc Good/no Good? Why?
Notes: Dark discs may show complete fusion at the surface but amy have inclusion inside. Shine a light through itto insure complete fusion.

## Slide 49
Dark Disk-Illuminate it!

## Slide 50
Recrystallization
Notes: Recystallization is usually due to small pieces of unfused material which initiates the process. Don't try to analyze the specimen. It may crack in the machine. It also is not a complete fusion and will bais the results.

## Slide 51
Bubbles in Analytical Surface- incomplete ignition?
Notes: Small bubbles usually occur in new molds for the first few fusions. Improperly prepared powders is another cause. Fineness is important. A completely ignited sample, that is ground fine enough is all that is required.

## Slide 52
Incomplete fusion. Not ground fine enough? / Wrong Flux?/Too cool when casting.

## Slide 53
Incomplete Disc/recrystallized - too much anti-wetting agent/wrong flux

## Slide 54
Care & Cleaning
Notes: Platinumware is not cheap. Use appropriate care. A dliute solution of HCl (5%) will clean the crucibles and moulds. NEVER use HNO3&HCl. Aqua regia will yield you a very nice solution of chloroplatinic acid. The crucible and molds will be so clean you will be able to see right through them (if you can see them at all)! Nothing better to use than invisible crucibles and molds.

## Slide 55
Touch-ups
Notes: When molds are mishandled the discs may show convex or concave surfaces. Rememeber that Io falsl off as the square of the distance between the sample and the tube. Low or high totals are usually the result. Lapping the surface flat using 30 micron diamond wheel may be needed.

## Slide 56: Liquids in XRF Analysis

## Slide 57: Questions
How many people have run XRF liquid analysis in the past week?
How many have the helium option?
How many need to routinely analyze liquid samples using other instrumentation, ie AA, ICP, IC, or traditional gravimetric, colorimetric , titrimetric methods?

## Slide 58: Thin film transmittance

## Slide 59: Film Degradation Resistance
Chemplex

## Slide 60: Cup, film holder

## Slide 61: Variety of films

## Slide 62: Film supports

## Slide 63: Size matters

## Slide 64: Cups, various sizes to fit XRF

## Slide 65: Liquid safety

## Slide 66: Bromine
Chart: 
| 100 | 10 | 1 | 0
Bromine | 117.9103 | 57.0623 | 50.7129 | 49.7801

## Slide 67: Chlorine
Chart: 
| 10000 | 100 | 10 | 1 | 0
 | 38.4965 | 0.8211 | 0.4831 | 0.4485 | 0.4291

## Slide 68: Sulfur
Chart: 
| 8015 | 33.39 | 3.339 | 0
 | 52.9081 | 0.4076 | 0.1749 | 0.1404

## Slide 69: Calcium
Chart: 
| 100 | 10 | 1 | 0
 | 1.9146 | 0.7629 | 0.6806 | 0.6559

## Slide 70: Magnesium
Chart: 
| 6000 | 100 | 10 | 1 | 0
 | 10.7392 | 0.5423 | 0.3243 | 0.3198 | 0.306

## Slide 71: Potassium
Chart: 
| 1000 | 100 | 10 | 1 | 0
Potassium | 12.4232 | 1.626 | 0.5998 | 0.5018 | 0.4924

## Slide 72: Sodium
Chart: 
| 6478 | 100 | 10 | 1 | 0
Sodium | 2.9126 | 0.1442 | 0.1016 | 0.1037 | 0.0993

## Slide 73: Water –Soluble Cr
Chart: 
| 100 | 80 | 40 | 4 | 0
 | 6.5725 | 5.4029 | 3.5957 | 1.9739 | 1.7843

## Slide 74: Inter-element Corrections
And now the math!
Didn’t you always want to know what happens for calculations!

## Slide 75: General form of influence correction is:
Ci (corrected) = Ci (apparent)(1+{Ʃ αij x cj}) Ci = concentration of element of interest aij = influence coefficient of element j on element i
j = refers to interfering element
    cj = refers to concentration of interfering element

## Slide 76: Iterative Calculations

## Slide 77: Sources of errors

## Slide 78: Errors due to counting statistics
  There will be a random error associated with the measured value of N Approximately 68% of all of the data lies within one standard deviation of the mean.
  Approximately 95% of all the data is within two standard deviations of the mean.
  Approximately 99.7% of the data is within three standard deviations of the mean.
Do not associate errors of counting statistics with error in concentrations

## Slide 79: Common errors associated with calibration
Failure to adequately separate instrumental and matrix dependent effects.
Poor judgement on the part of the analyst as to whether or not a correction should be made.
Poor technique on the part of the analyst in the determination of influence corrections. i.e. 2n2+1=Number of stds needed to apply. (n= number of inter-element correction’s)
Poor quality and/or range of calibration standards.
Inadequacy of the regression program.
Application of the technique where the specimens are inadequately homogeneous.

## Slide 80: Questions?

## Slide 81: Thank You
Don Broton
CTLGroup
5400 Old Orchard Rd. Skokie, IL 60077
DBroton@CTLGroup.com