## Slide 1: Cements Materials Characterization Workshop
Specimen Preparation for X-Ray Powder Diffraction
Phase Identification for Cementitious Materials

## Slide 2
No detailed procedure, providing flexibility
Qualification approach based upon each user’s analysis of SRM clinkers
Precision and accuracy criteria based upon an inter-laboratory study
Four cements compounded using SRM clinkers and lab-prepared calcium sulfates
Rietveld and traditional methods are acceptable
Some guidance provided for phase identification and analytical procedures
ASTM C1365

## Slide 3: Errors in Powder Diffraction
Relative Importance of Errors
scale d- Measurement Errors I – Instrument Errors Type
10 Displacement S
9
8
7 Preferred Orientation R
6 d-spacing R
d-reporting R
Stress/Strain R
5 Axial Divergence S
Zero Angle S
I – Reporting R
4 Flat Specimen S
Transparency S
2:1 2:1 S
3 Spec.-Rec Slit Distance S
Broad Lines S
Counting Statistics Counting Statistics R
2 Beam Uniformity R
alpha-2 influence
cocked slits
mechanical
1
0 Temperature Variations
wavelength CuKalpha
---------------------------------------------------------------------------------------------------------------
S = Systematic Error R = Random Error
Reproduced from:
W.N. Schreiner, C. Surdukowski, R. Jenkins
and C. Villamizar, “Systematic and Random Powder Diffractometer Errors Relevant to Phase Identification”,
Norelco Reporter Vol. 29, No. 1, April 1982
Specimen displacement has little effect on intensities
Preferred orientation continues to present problems, particularly with the top-loaded specimens commonly used today,
After displacement, round off error in d-spacing reporting is most significant
Alignment errors may be identified and corrected via external standards (zero error, 2:1)
Zero error and specimen displacement are highly correlated over short 2-theta distance
Displacement and orientation errors will affect phase identifications
Incorrect phase identification introduces bias
These focus on data collection, analysis is another aspect of the test procedure
Careful phase identification and refinement procedure improves the possibility of a success

## Slide 4: Preferred Orientation: Alite
29.3
30.07
32.1
32.7
34.2
Ground to – 10 µm
Specimen pressed to induce
preferred orientation (red trace)
Minimal
preferred orientation
Peak Intensity Ratio
29.3/30.1 32.1/32.7
2.6 1.2 OK
2.6 1.5 not OK
Coarse-ground samples, like cement from the finish mill, can be susceptible to preferred orientation.
Phases that exhibit strong cleveage like alite, calcite, gypsum are prone to orientation, particularly with top-loaded specimens.
For alite this is most apparent with the diffraction peak occurring at 32.1 degrees.
Fine-grinding, careful specimen loading and pressing minimize these tendencies, though for some phases with strong cleavage we rarely eliminate orientation
Orientation corrections are possible but seem to work best in quantitative analysis when orientation is not too severe.

## Slide 5: Specimen Preparation
Cement: As-Received
Cement: hand-ground, mortar and pestle
Cement: 70% <10 um
Modern instruments today obtain results in a fraction of the time required in 1930 but are still subject to sample preparation and mounting artifacts that may affect the analysis,
Grinding reduces particle size, improves representative sampling and particle statistics, reduces preferred orientation tendencies, and improves packing characteristics,
The packed powder mount ideal characteristics present a representative sampling of bulk material, a homogeneous volume, and smooth surface.
Cement Powder Specimens
2 mm field width
Cement phase solid interaction volume is roughly 15 µm at 40 degrees 2-theta
Sampling is improved by specimen rotation, bringing more particles into the beam footprint
Many particles are multi-phase with individual crystals smaller than the particle

## Slide 6: Preparing Powders
to -10 um
Jaw Crusher – down to mm-sized fragments
Mortar and Pestle – to -50 um
Splitter for subsampling

## Slide 7: Preparing Specimen Mounts
A wide range of specimen mounts are available for normal and specialized applications

## Slide 8: Preparing Powder Mounts
Pour approximate volume, level
Gently press with slide and paper
Examine surface for smoothness and uniformity
Spatula
Microscope slide
Cavity holder
Weighing paper
Determine the approximate amount to make a compact specimen mount. For these holders, about 1.2 g is sufficient.
Level the surface by cutting with the spatula and gently tapping the mount, press it flush using the weighing paper and glass slide.
Powder on the holder surface will result in height displacement!

## Slide 9: Clinker Phases
Clinker is comprised of a three-dimensional framework of alite and belite crystals and a matrix of aluminate and ferrite, crystallized from a molten liquid, periclase dispersed throughout as equant to dendritic crystals, alkali sulfates, which are typically located along fractures and pore walls. Finally, porosity may also be considered a component of the microstructure.
The phase texture, size and distribution reflect the history of the raw materials preparation, firing (clinkering), and cooling.
HFW Taylor, Cement Chemistry
DH Campbell, Microscopical Examination and Interpretation of Portland Cement and Clinker

## Slide 10: Silicate Phases
Alite (Ca3SiO5): comprises 40 % to 70 % of a clinker, has triclinic, and monoclinic forms; impurities of up to about 4% stabilizes the monoclinic form; magnesium and aluminum may substitute for silicon and other substitutions include iron and sodium; distinct (idiomorphic) crystals exhibiting six-sides; density may range from 3.13 g/cm3 to 3.22 g/cm3
Belite (Ca2SiO4): comprises 10 % to 30 % of a clinker as a series of orthorhombic and monoclinic solid solution phases; impurities of up to 6 % as Mg, K, Na, Al, Mn, P, Fe and S may stabilize the polymorphs such as the beta form (most common), alpha, alpha-prime, gamma; as idiomorphous, rounded grains often exhibiting a lameller texture due to twinning; density (beta form) is 3.28 g/cm3

## Slide 11: Interstitial Phases
Aluminate (Ca3Al2O6): comprises up to 18 % of a clinker as xenomorphic or anhedral (no distinct form) crystals intermixed with the ferrite phase has cubic, orthorhombic, and monoclinic forms; impurities of up to about 10 % include Mg, Si, K, and Na; alkali aluminate is typically orthorhombic and monoclinic, may exhibit a lath-like texture; density around 3.064 g/cm3 for the cubic form and 3.052 g/cm3 for orthorhombic
Ferrite (Ca4Fe4O10 to Ca4FeAl3O10): is an orthorhombic solid solution comprises about 14 % of a clinker; impurities of up to 13 % as Mg, K, Na, Al, Mn, and Ti; exhibiting a lath-like (slow cooling) to dendritic (rapid cooling) texture; density is 3.77 g/cm3

## Slide 12: Dispersed Phases
Free Lime (CaO): Cubic, rounded, idiomorphous grains, sometimes agglomerated and retaining the shape of the parent limestone fragment; density of 3.32 g/cm3
Periclase (MgO): Cubic, euhedral (equiaxed),rounded to dendritic grains free within the clinker and as inclusions in the silicates and interstitial phases, comprising up to 7 % of a clinker; density of 3.58 g/cm3
Na and K sulfates may only comprise up to a few percent of the clinker but can affect early hydration as they typically form along pore and fracture surfaces and have a high solubility.
Arcanite (K2SO4); orthorhombic; density of 2.67 g/cm3
Aphthitalite ((K2Na)2SO4); trigonal; density of 2.66 to 2.71 g/cm3
Ca Langbeinite (K2SO4 2CaSO4); cubic; density of 2.87 g/cm3 Thenardite (Na2SO4) ; orthorhombic; density of 2.67 g/cm3
Syngenite (K2Ca(SO4)2 H2O; monoclinic, secondary product in cements density of 2.6 g/cm3
BE-K-S

## Slide 13: Phase Identification: 1
Traditional methods - manual search-match using the Hanawalt method.
Identify significant peak positions and relative intensities
  correction for displacement errors would requite some kind of internal standard.
  otherwise, keep in mind the need for some flexibility in peak locations
  relative intensity is useful but secondary to peak position and the collective set of peaks
List peak d&I in order of relative intensity

## Slide 14: Phase Identification: 1
Traditional methods - manual search-match using the Hanawalt method.
List peak d&I in order of relative intensity
Search on strongest peak followed by second and third strongest.
Each entry is listed three times to account for relative intensity error
Any guesses?

## Slide 15: Computer-Aided Search-Match
Mark peak positions and relative intensities
Set negative and positive chemistry
set subfiles; user-created subfile if possible

## Slide 16: Computer-Aided Search-Match
Run the search

## Slide 17: Computer-Aided Search-Match
No chemical restraints or subfile selection can lead to MANY hits

## Slide 18: Computer-Aided Search-Match
Check the stick figures against the pattern; consider the plausibility of the chemistry
Some are close!

## Slide 19: Computer-Aided Search-Match
Check the stick figures against the pattern; consider the plausibility of the chemistry
Some are very close!

## Slide 20
Check against the ICDD
Database Card Files using
either the D & I values or
a graphical representation
as a stick figure against the
diffraction pattern.
Recall that the specimen
displacement error is the
greatest so expect some peak
shifting.
For clinker phases, chemical
substitution will also affect
the peak positions.
The high degree of pattern
overlap and the number of
phases also complicates
identification. Therefore, an
alternative is to look for the key
resolvable peaks and then check
the cards.
Peaks not assigned to a phase
may be from some new phase

## Slide 21: Key Lines for Cement Phase Identification using XRD

## Slide 22: Key Lines for Cement Phase Identification using XRD

## Slide 23: Practicum 2
Using the file Cement Key Lines.pdf, identify the single phases and simple mixture phases found in the folder XRD Phase Identification
Each folder contains
an image of the pattern (.JPG),
a .RAW powder diffraction data file (.RAW) that may be read into Profex,
and a text file providing peak d-spacings, two-theta positions, and relative intensities
Work on identifying the file using the key lines and knowledge that it diffraction pattern is from a phase that is related to portland cement
Work in pairs if you wish
While most examples are of a single phase, one or more may be a mixture

## Slide 24: Analyis of Clinker and Cements Using Selective Extractions

## Slide 25: Phase Identification
ASTM C1365: Standard Test Method for Determination of the Proportion of Phases in Portland Cement and Portland-Cement Clinker Using X-ray Powder Diffraction, Annual Book of ASTM Standards, ASTM Int., West Conshohocken, PA, 2013.
Cement phases are a challenge and can be difficult to identify using the traditional methods
Tentative identification is made using diagnostic peaks for each phase rather than the traditional methods (three major peaks); confirmation uses the ICDD powder diffraction database
g
f
g
b
B
A
a
Al
f
Table X1.1 of ASTM C1365 lists peaks by descending d-spacing that have minimal interference. ICDD card numbers allow recalling the entire pattern.
b
g – gypsum
f – ferrite
b – bassanite
a – anhydrite
A – alite
B – belite
Al – aluminate

## Slide 26
SRM 2686a
Bulk,
KOSH: silicates, periclase
SAM: ferrite, aluminates, alkali sulfates, periclase, (calcite and calcium sulfates for cements)

## Slide 27: Selective Extractions
Bulk Clinker and KOSH Residue (silicates, periclase)
Bulk Clinker and SAM Residue
Aluminates, ferrite, periclase, alkali sulfates, and for cements, calcium sulfates
Provide a clearer picture of phase groups to facilitate identification, refinement of their characteristics (peak shape, lattice parameters), and an additional sample set for quantitative analysis.
1) KOH-Sucrose: 7.5 g KOH, 7.5 g sucrose, 75 ml water. Heat to 95 C, stir in 2 g ground cement for one minute. Vacuum filter, wash residue with 20 ml water followed by 50 ml of methanol
2) Salicylic Acid / Methanol: 20 g salicylic acid in 300 ml methanol. Stir 5 g ground cement in SAM solution for 2 h, vacuum filter, wash residue with methanol, dry below 60 C, weigh residue
3) Nitric Acid / Methanol: 10 g of ground clinker or cement in 500 ml of 7 % nitric acid in methanol for 30 minutes to produce residue of ferrite and periclase (if present).

## Slide 28: Analytical Approach
Bulk Analysis
  detection limits due to dilution by silicates make them difficult to identify – but certain diagnostic peaks are evident even in low concentrations. An alternative is to perform a selective extraction to concentrate the intersitital phases, calcium sulfates, and alkali sulfates
SAM extraction residue
  quantitative extraction, QXRD and re-calculation of phases on a whole-cement basis
KOSH Extraction Residue
  can be a difficult extraction but will provide an improved pattern of the silicate phases and identification of belite and alite forms
  Generally not quantitative but still useful

## Slide 29: Common Sulfate Phases in Clinkers and Cements
arcanite (K2SO4) may accommodate Na+, Ca2+, and CO3 in solid solution
aphthitolite (K4-x,Nax)SO4 with x usually 1 but up to 3)
calcium langbeinite (K2Ca2[SO4]3) may occur in clinkers high in K2O and SO3
anhydrite (rare) (CaSO4) in clinkers with high SO3 (>2%)
thenardite: (rare) (Na2SO4) in clinkers with high Na/K ratios
these phases are generally found along silicate crystal boundaries – though sometimes too as inclusions
Some sulfate is substituted in the silicates (below 1% for low-sulfate and 2% for high-sulfate clinkers. Some may also be found in the aluminates and ferrite.
SO3 to alkali ratio of up to 0.5 occurs as alkali sulfates. Higher SO3 levels, there is insufficient alkali dissolved to balance the sulfate and that sulfate may be present in other phases – calcium langbeinite, anhydrite, silicates

## Slide 30: Cement Sulfates
gypsum: CaSO4 2[H2O]
bassanite: (hemihydrate, plaster) 2CaSO4 H2O, or γ-CaSO4
anhydrite: CaSO4
syngenite: (K2Ca2SO4 [H2O]) formed by reaction of gypsum and potassium sulfate - reacts rapidly on mixing.

## Slide 31: Calcium Sulfates
Gypsum (CaSO4 2H2O) – moderate solubility. Dehydrates around 100 C to bassanite (hemihydrate or plaster) CaSO4 ½ H2O and at higher temperatures to anhydrite (CaSO4). Both dehydration products exhibit different solubility than gypsum
Bassanite hydration to gypsum during mixing may result in paste stiffening – false set
The SAM extraction enhances detection limits by eliminating the dilution effect of the calcium silicates. Be careful however, to dry the residue at low enough temperature that dehydration does not alter the forms

## Slide 32
Aluminate Forms: Cubic and Orthorhombic
cubic – two peaks around 21 and 21.8 degrees;
single peak at 33.2 degrees. The 21 degree peak
has half the intensity of the 21.8 degree peak
orthorhombic: the 33.2 degree peak is split and there
is no peak at 21.8 degrees.
There is an overlap with the 21 degree cubic form peak.
If the 21 degree peak is greater than half the intensity of
the 21.8 degree peak, you may conclude that there is some
orthorhombic aluminate in the sample
This is best observed in a SAM extraction residue

## Slide 33
triclinic vs. monoclinic alites
the 51 degree region is
relatively clear of other
peaks and is a good place
to look – especially in the
KOSH residue

## Slide 34: Calcite in Cement
Peak Intensity Ratio
29.3/30.1: 2.6
alite
calcite

## Slide 35: Composite Pattern 2686a

## Slide 36: Alite

## Slide 37: Belite

## Slide 38: Aluminate: Cub., Ort.

## Slide 39: Ferrite

## Slide 40: Periclase

## Slide 41
CCRL 142 Data Set
KOSH - 76.04% of cement, SAM - 23.96% of Cement
BULK KOSH SAM
C3S 54.5 48.9
b-C2S 21.6 18.1
a-C2S -- 0.9
C4AF 10.1 10.3
C3Ac 5.0 4.2
Arc. 0.4 0.1
Lang. 1.3 1.3
Peric. 3.2 3.5 3.7
Anh. 0.4 1.2
Bass. 2.8 4.1
Gyp. 0.1 0.2
Qtz. 0.1 0.2 0.1
Dolo. 0.8 1.0
Recalculation of Selective Extraction Data
- provides an extra estimate based upon a concentrated residue of the selective extraction
IR
(43)
(17.5)
(0.4)
(5.4)
(15.4)
(5.0)
(17.1)
(0.8)
(0.4)
(4.2)

## Slide 42: Practicum 3: SRM2686a
Start with phase identifications using the selective extraction patterns in the IMAGES folder
Examine images SRM2686a_KOSH_d&I_1.png and SRM2686a_SAMd&I_1.png
and their corresponding .txt files
With the knowledge that you are working with a portland cement clinker, look for key diffraction peaks for phases you might encounter in that material
Prepare a list of potential phases
Starting with the powder diffraction files found in the Bulk folder, read one of the KOSH or SAM patterns into profex, load the appropriate instrument file from the Device folder (D8_6Div_4SS.sav), load the appropriate phase files from the CementStructures folder and initiate the refinement. Transfer the results to the spreadsheet
Repeat for each of the two or three replicates for the SAM and KOSH
Proceed to the Bulk data files (86a_10_B_0*.raw) and repeat