## Slide 1: Cements Materials Characterization Workshop
July 12 – 14, 2017
Paul Stutzman
Jeffrey Bullard
Steve Feldman
Alexander Brand
LaKesha Perry
Don Broton
-C2S

## Slide 2
Lecture Rooms A & B
Lecture Room A
Lecture Room B
To 101 Parking Lot
Main Lobby
OUTSIDE EXIT
OUTSIDE EXIT
Stairs to Basement Exits and Shelter in Place
Shelter in Place
In Room

## Slide 3: NIST
Bldg. 101

## Slide 4: Workshop Configuration
We will focus on a number of materials characterization applications: X-ray powder diffraction (XRD), scanning electron microscopy (SEM) with X-ray imaging (XR), image processing and image analysis (IA) , X-Ray fluorescence (XRF), and Digital holographic microscopy (DHM).
The XRD and SEM portions have an introductory presentation followed by a live demonstration and a lab practical
The exercises have real lab data for the NIST SRM clinkers and portland cement
The SRM certificates are provided for the clinkers for comparison of your estimates, and for the XRD exercise, the ASTM C1365 qualification criteria are provided

## Slide 5: Early Work: Clinker Manufacture
Aspden Patent 1824
from The Chemistry of Portland Cement, RH Bogue, 1954
It had been recognized that some limes were hydraulic and some were not but the reasons remained a mystery. John Smeaton found that the best limes were impure, containing a substantial amount of clay (25%), and additionally used a pozzolanic tuff. (1756)
Similar findings appear to have been made independently in other countries
The reasons the materials were hydraulic remained a mystery
Bottle kilns at Lehigh PA
D. Campbell, CTL

## Slide 6: Le Châtelier and Tornebohm – the first look
Microscopy of thin sections, 1883
description of the clinker phases
postulate compositions
observed decomposition of SiO2.2CaO
relating excess free lime to durability problems
1905, translation by J.L. Mack
https://books.google.com/books/about/Experimental_Researches_on_the_Constitut.html?id=PVFDAAAAIAAJ
500um
crossed polars SRM2688
Plane-Polarized light: SRM2688

## Slide 7: Geophysical Laboratory, Early 1900’s
A.L. Day, E.S. Shephard, F.E. Wright “The Lime Silica Series of Minerals”, Am. Jour. Sci. (4) xxii, 265 1906
E.S. Shepherd, G.A. Rankin, F.E. Wright, The Binary Systems of Alumina, with Silica, Lime, and Magnesia, Am. Jour. Sci. (4) xxviii, 293, 1909
E.S. Shepherd, G.A. Rankin, F.E. Wright, “Preliminary Report on the Ternary System CaO-Al2)3-SiO2. A Study on the Constitution of Portland Cement Clinker”, J. Ind. Eng. Chem., iii, 211-227, 1911
Light microscopy was limited by the equipment of the day, the subtle differences in optical properties, and the fineness of the constituent phases but remained useful for examination of laboratory-prepared mixtures
Three general theories prevailed on cement:
  a sold solution of various silicates and calcium aluminates
  a definite lime, silica and alumina compound
  composed of separate silicates and aluminates, some or all of which are hydraulic
Work at the Geophysical Laboratory on binary systems CaO-Al2O3, CaO-SiO2, and Al2O3-SiO2 and the ternary system CaO-SiO2-Al2O3 using many compounded samples and optical microscopy concluded that cement is composed of separate silicates and aluminates

## Slide 8
Thin section of dicalcium silicate, crossed Nichols, 90x
Tricalcium silicate (gray), along with dicalcium silicate crossed Nichols, 190x
“At early periods the constituents of Portland cement of normal composition and manufacture, in the order of their strength-conferring properties, are: tricalcium silicate, tricalcium aluminate, and dicalcium silicate,
At periods beyond 28 days the dicalcium silicate gains sufficient strength to place it almost on equality with the tricalcium silicate,
Tricalcium aluminate containing 10 percent plaster gains practically no strength after the first period at which it was tested; that is, 24 hours,
Tricalcium silicate of the purity used in this investigation (90 per cent 3 CaO.SiO2 in one case and 95 per cent in the other) has all the important properties of Portland Cement, especially those of the “rate of setting” and strength developed,
dicalcium silicate, such as used in this investigation, sets too slowly and attains strength too slowly to be of any commercial value when used alone,
Tricalcium aluminate alone, as used in this investigation, sets too rapidly and attains too little strength to be of any commercial value as a hydraulic cementing material
Plaster of Paris, when added to any of the compounds or mixtures of those studied, generally increased their strength. This effect is more marked at the early periods.”
P.H. Bates and A.A. Klein, Properties of the Calcium Silicates and Calcium Aluminate Occurring in Normal Portland Cement, Technologic Papers of the Bureau of Standards, No. 78, 1917

## Slide 9: Cement Chemist Nomenclature
Cement chemists use a variety of methods for expressing chemical formulae; 1) standard, 2) a sum of oxides and, 3) an abbreviated oxide nomenclature that stems from the Geophysical Laboratory work of Rankin and Wright (1915).
For example, tricalcium silicate (Ca3SiO5) may also be expressed as 3CaOSiO2 and as C3S
Simple oxide phases such as free lime (CaO) or periclase (MgO) generally are expressed without abbreviation
C = CaO S=SiO2 A = Al2O3 F = Fe2O3
M = MgO K = K2O S=SO3 N = Na2O
T = TiO2 P = P2O5 H = H2O C = CO2
To ensure things remain confusing, C3S or tricalcium silicate usually refers to the pure phase (i.e., no chemical substitution) while alite refers to the impure industrial products. This convention is often not followed and the names are used inter-changably

## Slide 10
1930 X-Ray Powder Diffraction Camera work by the Portland Cement Association Fellowship at the National Bureau of Standards
At that time there were three prevailing theories on phase composition of portland cement clinker:
  a complex compound containing lime, alumina, and silica
  lime and silicates form a loosely bound solid solution
  lime and silica form two separate silicates
Comparison of synthesized phases to commercial clinkers demonstrated the presence of tricalcium silicate as the primary phase, as postulated by LeChatelier previously (1882, 1905), observed by Tornebohm (1887) independently, and the work of Rankin and Wright (1915) on the system CaO-Al2O3-SiO2
Other phases identified included what are now called belite, aluminate, ferrite and periclase.
X-ray diffraction powder camera patterns required 50 h exposure times!

## Slide 11
1930 Powder Camera work by the Portland Cement Association Fellowship at the National Bureau of Standards
Museum of Optical Instruments
Debye-Scherrer Camera
Philips Technical Review, Vol. 16, No. 4, 1954
The wide-angle goniometer introduced in the late 1940’s provided improved resolution, more stable and intense X-ray sources and an output to a strip chart that facilitates position and intensity measurement
Quartz, SiO2
Cristobalite, SiO2
Glass, SiO2

## Slide 12: Cements-Related Developments in XRD
PCA Fellowship: Bogue, Brownmiller, Harrington, Ordway
L.E. Copeland and R.H. Bragg: Quantitative analysis procedures with or without calibration curves, including a means to deal with overlapping lines (1958)
PCA controlled atmosphere (RH, N2) experiments on cement hydration
G. Frohnsdorff, R. Berger, P. Harris, P. Johnson: automated powder diffractometer (1964), sample changer and least-squares decomposition of diffraction patterns
PK Mehta, M.J. Shah: Standards-based XRD analysis of clinker for closed-loop kiln control
Studies on precision by AE Moore (1964), L. Aldridge (1982)
1978: XRD Task Group started in ASTM C01.23, Compositional Analysis
  initial data using sets of pure phases and calibration curves
  development of the NIST Reference Clinkers to develop and test the methods
  development of the first ASTM standard C1365 for XRD analysis of cements
  amending the standard XRD test to include Rietveld analysis

## Slide 13: ASTM C01 Cement Committee
Develops specifications, methods of test, recommended practices,
and terminology for hydraulic cements
C01.23 Subcommittee on Compositional Analysis
  To develop and maintain test methods and practices for determination of chemical and mineralogical composition of hydraulic cements
  XRD Task Group established in 1978
C114 Standard Test Methods for Chemical Analysis of Hydraulic Cement
C1356 Standard Test Method for Quantitative Determination of Phases in Portland Cement Clinker by Microscopical Point-Count Procedure
C1365 Standard Test Method for Determination of the Proportion of Phases in Portland Cement and Portland-Cement Clinker Using X-Ray Powder Diffraction Analysis

## Slide 14: ASTM C01.23 Interstitial Phase Calibration
Early efforts used individual peak area measurement and calibration curves plotted as mass ratios vs. intensity ratios
Pattern-fitting improved the precision of the estimates of the calibration curves and for the unknown specimens
The first version of ASTM C1365 covered the interstitial phases based upon calibration curves developed using mixtures of pure phases
Compensating Polar Planimeter
Pattern-fitting for QXRD of portland cement clinker, ICMA 1996

## Slide 15
Periclase
C3A-O
C3A-C
Ferrite
C2S-alpha
C2S-beta
C3S
C3S - mono
Figure 4
best-fit pattern (green)
difference plot

## Slide 16: Phase Abundance Analysis
Simultaneous refinement of X-ray diffraction patterns of multiple phases allows quantitative analysis:
Wp = (Sp (ZMV)p) / (Σi[Si (ZMV)i])
where
Wp - the mass fraction of phase p
S - the Rietveld scale factor
Z - the number of formula units per unit cell
M - the mass of the formula unit
V - the unit cell volume

## Slide 17: SRM Clinkers
SRM 2686a: Heterogeneous texture with alite, clustering of belite, a differentiated matrix comprised mostly of ferrite and some fine-grained aluminate, and uniform distribution of periclase.
SRM 2687: Fine-grained clinker with abundant alite, some small nests of belite, variable porosity, and abundant, undifferentiated matrix, consisting of both a medium and fine-grained ferrite
SRM 2688: Coarse-grained clinker with uniform phase distribution, a medium- to coarse-grained well differentiated matrix, and occasional presence of periclase and alkali sulfates within the matrix and along grain boundaries.

## Slide 18: Uncertainty in Bogue-Calculated Phase Content of Portland Cement
The traditional means of estimating potential cement composition using the Bogue calculations (1929) has limits recognized since inception.
oxide uncertainty
1- mass percent
oxide and constants uncertainty
1- mass percent2
Bogue Calculations
1 P.E. Stutzman, and D.S. Lane, “Effects of Analytical Precision on Bogue Calculations of Potential Portland Cement Composition,” Journal of ASTM International, Vol. 7, No. 6
2 P. Stutzman, A. Heckert, A. Tebbe, and S. Leigh, “Uncertainty in Bogue-Calculated Phase Compositions of Hydraulic Cements”, Cement and Concrete Research, 61 (2014) 40 – 48
1
Uncertainties in bulk oxide determinations,
Uncertainties in the Bogue constants, reflecting deviations of clinker phase chemistry from ideal assumptions,
Not accounting for all phases, and
Covariances among oxide errors,
Alite

## Slide 19: SRM Clinker Certificate

## Slide 20: ASTM C 150 Limits
Currently, these phase mass fractions are estimated using the Bogue calculations of potential composition from a bulk chemical analysis.
We can relate the existing specification criteria to XRD-derived phase measurements via calibration

## Slide 21: Clinker Microscopy
Useful clinker microscopy texts:
D.H. Campbell (1999) Microscopical Examination and Interpretation of Portland Cement and Clinker, 2nd ed., Portland Cement Association, Skokie/IL
F. Hofmänner (1975) Microstructure of Portland Cement Clinker, Holderbank Management and Consulting, Ltd., Holderbank/CH
H. Insley and V. Fréchette, (1955) Microscopy of Ceramics and Cements, Academic Press, NY
A useful chart for interpretation of clinker textural features is provided by F. Hofmänner (1975).
Using the image of SRM 2686 above, with medium-grained silicates exhibiting heterogeneous distribution, coarse matrix, free lime content little to none reflects a heterogeneous raw mix, normal burning conditions and slow cooling
While point count procedures (ASTM C1365) may be used to quantify features, visual estimates remain a very useful tool in evaluating the clinkering process.

## Slide 22
No detailed procedure; provides flexibility but increases potential for precision and bias
Qualification approach based upon each user’s analysis of SRM clinkers
Precision and accuracy criteria based upon an inter-laboratory study
Four cements compounded using SRM clinkers and lab-prepared calcium sulfates
Rietveld and traditional methods are acceptable
Some guidance provided for phase identification and analytical procedures
ASTM C1365

## Slide 23: Qualification
11.1.2 Prior to use for analysis each QXRD test procedure must be qualified. The laboratory shall maintain records that include a description of the procedure and the qualification data. These records shall be made available to the purchaser.
11.1.5 Qualification shall consist of analyzing the three SRM clinkers for the proportions of C3S, C2S, C3A (cubic and orthorhombic), C4AF, and M.
11.2.2 If replicate analyses are carried out, results should differ from each other by no more than the value shown in Table 1.
11.2.3 The mean result should differ from the known value by no more than the value shown in Table 2 for the number of replicates used.
The inter-laboratory study included the calcium sulfates and calcite in anticipation of the analysis of cement, but those phases are not specified in the C150 specification

## Slide 24
The numbers are pooled across four cements and are expressed in 1-s terms.
To convert sr and sR to ASTM r and R multiply by 1.96*√2
 | Repeatability Within-lab |  | Reproducibility Between-lab | 
 | s within | r | s between | R
Alite | 0.74 | 2.04 | 2.23 | 6.18
Belite | 0.64 | 1.77 | 1.41 | 3.91
Aluminate | 0.47 | 1.31 | 0.74 | 2.05
Ferrite | 0.49 | 1.36 | 0.95 | 2.63
Periclase | 0.23 | 0.63 | 0.32 | 0.89
Arcanite | 0.22 | 0.60 | 0.41 | 1.13
Gypsum | 0.21 | 0.59 | 0.58 | 1.62
Bassanite | 0.39 | 1.08 | 0.81 | 2.24
Anhydrite | 0.27 | 0.74 | 0.63 | 1.75
Calcite | 0.99 | 2.73 | 0.50 | 1.50
The within-laboratory standard deviation for each phase is reported as ‘s-within’.
Therefore, results of two properly conducted tests by the same operator should not vary more than ‘r’.
The multi-laboratory standard deviations are ‘s-between’. Therefore, results of two properly conducted tests on the same clinker or cement by two different laboratories should not differ from each other by more than ‘R’.
Notes: Pooled BOB data for repeatability, maximum likelyhood data for between

## Slide 25: Table 2: Prediction Intervals
Phase | 2 replicates | 3 replicates | 4 replicates
Alite | 5.93 | 4.91 | 4.31
Belite | 3.70 | 3.06 | 2.69
Aluminate | 2.14 | 1.77 | 1.55
Ferrite | 2.46 | 2.04 | 1.79
Periclase | 0.77 | 0.64 | 0.56
Arcanite | 0.85 | 0.70 | 0.61
Gypsum | 1.55 | 1.28 | 1.12
Bassanite | 1.52 | 1.26 | 1.11
Anhydrite | 1.67 | 1.38 | 1.21
Calcite | 0.68 | 0.56 | 0.49
Accuracy:
“the closeness of agreement between a test result and a accepted reference value”,
- includes both random and systematic error
The mean result shall differ from the known value by no more than the value shown in Table 2 for the number of replicates used.
The prediction intervals (95%) are designed to bracket values of a mean of k (=2,3,4) future measurements of the relevant phases

## Slide 26: Cements Analysis

## Slide 27: Establishing equivalent QXRD values for C150 limits
P. Stutzman and S. Leigh ” Statistical Calibration of ASTM C150 Bogue-Derived Phase Limits to Directly Determined Phases by Quantitative X-Ray Powder Diffraction,” Jour. ASTM International Vol. 7, No. 7, July 2010
Alite
Belite
Aluminate
Ferrite

## Slide 28: Calibration: C150/QXRD
Equivalent XRD values based upon existing specification limits that were developed from historical performance data

## Slide 29
SRM 2686a Box Plots by XRD (left) and SEM
SRM certified phase values represent a consensus between XRD and SEM/X-ray/image analysis
The closeness of agreement for relatively low (periclase) and high (ferrite) absorbing phases suggests that microabsorption is not a large issue, at least for powders ground to 99% below 10 µm
Mass Attenuation Coefficients
Phase µ /p (g/cm2)
Alite 102
Belite 94
Aluminate 86 – 90
Ferrite 129 (underestimate potential)
Periclase 28 (overestimate potential)
Arcanite 87
Clinker 92
ferrite
periclase

## Slide 30: Cements Inter-Laboratory Study
  ASTM requirements for a standard test method include a precision and bias statement
  This it accomplished through an inter-laboratory study
  Results of the study may also provide consensus on recommended procedures
Four cements prepared using NIST SRM clinkers with addition of known amounts of laboratory-prepared calcium sulfates and calcite
  gypsum, bassanite, anhydrite
Triplicate analyses with sample re-packing
Qualitative analysis, reporting results and procedures
Statistical analysis to determine within- and between-laboratory precision and bias

## Slide 31: Precision
An assessment of the variability one may expect when the test method is used by one or more reasonably competent laboratories expressed as a standard deviation (1-σ)
Repeatability: within-laboratory variability
Reproducibility: between-laboratory variability
ASTM repeatability (r) and reproducibility (R) are derived by multiplying the appropriate standard deviation by the factor 1.96*√2, representing a 95% coverage factor
Bias
= (Estimate of Truth) – (Truth)
= (Estimate of True Mean Phase Conc.) – (Best Estimate of True Mean)

## Slide 32
known
(alite)
known
(belite)
Lab Values, Consensus Means, Standard Deviations, Known Values: Mixture #1
Note that for some labs (#11 for example), the constant sum error can affect all phase estimates. For others (#6), additional bias occurred for the interstitial phases. Both of these labs were deemed outliers and were excluded in precision calculations

## Slide 33
Precision: Youden Plot
The Youden scatter plot compares within- and between-laboratory variability from measurements from two cements, constituting x- and y- coordinates to form a point
The consensus value is based upon an 80 % trimmed mean of means. Boxes represent 1, 1.5, 2, and 2.5 standard deviations from the mean based upon the ASTM C1365 1 reproducibility precision, and are ranked from 4 to 0, respectively.
Points lying near the 45-degree reference line, but far from the median intersection reflect large systematic error
Points lying far from the 45-degree line indicate large random error
Accuracy
Precision

## Slide 34: Calcite
Precision: Youden Plot
The replication allowed calculation of a new set of precision and qualification criteria to update that currently in ths standard, but is based upon only one cement. These cements both contained a limestone addition.
Accuracy
Precision
Material | Sr | sR | r | R
A | 0.1936 | 0.9498 | 0.54 | 2.66
B | 0.2176 | 1.0517 | 0.61 | 2.94

## Slide 35: Light Microscopy